## Slide 1: Transducers: Electrodes for Biosensors
2. Polarization of electrodes:
(1) theoretically, half-cell potential E0 is defined for I=0 between the
electrode and electrolyte under chemical equilibrium.
(2) if I≠0, E0 will be altered, i.e., E0(I≠0) ≠ E0(I=0), due to polarization.
(3) the difference is defined as overpotential:
E0 = E0(I≠0) - E0(I=0).
(4) E0 is caused by the following factors:

## Slide 2: Transducers: Electrodes for Biosensors
2. Polarization of electrodes:
Ohmic:
Resistance R in the electrolyte causes a voltage drop along the path of
current I in the electrolyte (especially in low-concentration electrolyte).
It is I-dependent and does not necessarily follow ohm’s law.
Concentration:
If I≠0, oxidation and reduction reaction equilibriums are broken, i.e.,
resulting in a different E0.
Activation:
Any chemical reaction needs to overcome an energy barrier (activation
energy) governing the kinetics of the reaction to occur.
Oxidation CCn++ne- and reduction CCn++ne- do not necessarily
have the same energy barrier. If I is applied, either oxidation or reduction
reaction predominates; hence, the energy barrier is direction dependent.
The barrier difference at the interface appears as difference in potential
at the interface. This difference is activation potential btw solid/lqd phases.
,

## Slide 3: Electrodes: half-cell potential
Summary:
1. Electrolysis:
via electrochemical reactions to allow charge transfer across the interface.
2. Half-cell potential E0 depends on:
Type of material C of the electrode (e.g., E0);
[Am-], [Cn+];
T,
surface condition at the interface…
3. Why called half cell?
E0 is floated, therefore a second is needed to form a full cell to measure it.
V= E01 - E02.
4. There are two types of electrodes:
Potentiometry: by measuring voltage V (no need for polarization, I0);
Amperometry: by measuring current I (need polarization Eb, I0).

## Slide 4: Transducers: Electrodes for Biosensors
1. Polarizable & nonpolarizable electrodes
2. Measurement of E0
3. Nernst equation
4. Electrode equivalent circuits & skin model
5. ECG, EKG & EMG
6. Noise reduction in ECG sensing
Highlight

## Slide 5: Transducers: Electrodes for Biosensors
1. Polarizable electrodes:
The classification depends on what happens at the interface when an electrical current I is applied to the electrode.
Polarizable:
(Perfectly) polarizable electrode refers to those having no actual or very minute charge Q across the electrode-electrolyte interface, i.e., I0.
Definitely there will be I across the interface to form a close loop, but this I is a displacement current ID, behaving like a capacitor.
Most likely, those composed of noble metals, e.g., Pt, Au, which are very difficult to oxidize and dissolve.
ID relates to [Cn+] at the interface; therefore, it has a large concentration
related overpotential
Increase or decrease of charges at the interfaces results in
local ionic concentration changes

## Slide 6: Transducers: Electrodes for Biosensors
1. Nonpolarizable electrodes:
Nonpolarizable:
(Perfectly) nonpolarizable electrode refers to those having charge or current across the electrode-electrolyte interface freely, I0 requiring no energy to make the transition, i.e.,
A typical example is Ag/AgCl electrode, which will be discussed in more detail later.
Most likely, those composed of a metal C coated with a layer of slightly soluble ionic compound of same metal, Cn+ (in solid phase).
In these electrode, the electrolyte solution has to have a highly soluble anion and has to be saturated in order to be stable.

## Slide 7: Transducers: Electrodes for Biosensors
2 Measurement of E0:
More practical setup
complete reaction
2 half-cell reactions:
(battery)
(E01:=0.3V)
(E02:=- 0.76V)
(E0 = E01 - E02:=1.1V)
-SO4
-SO4
Zn2+
Cu2+
Suppose the internal resistance R of the full cell is 1kohm, then the current is 1.1mA.
However, depending on the type of electrodes (polarizable or not), R can be very
different, e.g., 10-1012ohm, thus the measurement in some cases can be challenging.
Keep in mind that these measurements are performed at 1M ion concentration. The
half cell potential in this case is defined as standard half-cell potential EØ.
Ion bridge!
Junction potential

## Slide 8: Transducers: Electrodes for Biosensors
2 Measurement of E0:
Measurement of E0 can be done with
a H2 electrode (E0=0).
M = Cu or Zn
A H2 electrode is an irreversible and almost non-
consumable electrode. It needs H2 bubbling, which is potentially explosive.
The measured potential
V = E1 - EH = E1 = 0.34V for M = Cu.
Similar measurements can be performed for most metallic electrodes by standard
H2 electrode. This is because most metals can be dissolved in an acidic solution.

## Slide 9: Transducers: Electrodes for Biosensors
Salt Bridge: contains an electrolyte that connects the 2 compartments of a voltaic cell,
allowing ion flow w/o extensive mixing of the different solutions, e.g., KCl.
mooni.fccj.org/~ethall/2046/ch18/galvanic.htm
The salt bridge is necessary to complete
the circuit and maintain charge neutrality
(stable and slow).
commons.wikimedia.org/wiki/Image:Electrochemi...
www.steve.gb.com/science/thermodynamics.html
?

## Slide 10: Transducers: Electrodes for Biosensors
2 Other 2 common electrodes:
Ag
AgCl
Hg
Ion bridge
Ag/AgCl:
Hg/Hg2Cl2:
Saturated (1M)
Water soluble with high Cl- solution, so they are compatible with biological uses
Calomel
electrode
Ag/AgCl provides
far better noise
performance than
Ag especially at
low f.

## Slide 11: Transducers: Electrodes for Biosensors
Ag/AgCl electrode:
(1) Composition: Ag: metal – nonsoluble;
AgCl: compound – slightly soluble;
Cl-: solution/electrolyte with high [Cl-].
(2) Reaction:
Ag  Ag+ + e-
Ag+ + Cl-  AgCl (eventually stabilized)
(3) Formation:
Ag is inserted into the Cl- solution with 5V, i.e., Ag Ag+ + e-;
because of high [Cl-], Ag+ + Cl-  AgCl, and a film of AgCl
is formed. Of course, they are other ways to make it.
depending on Ag doping, resistant R can be vastly changed.
(4) Static condition:
Solubility KS = aAg+ x aCl- =10-10 (because of AgCl). For most
biological applications, because of high [Cl-], e.g., [Cl-]0.1-1,
thus aAg+ 10-10 0, i.e., it is stable in a solution with stable [Cl-]. This is why it can be used for biological measurement.

## Slide 12: Transducers: Electrodes for Biosensors
3. Nernst equation in electrode-electrolyte reactions:
As one type of chemical reaction, electrolysis can be analyzed by Gibbs free energy:
G = -nFE (3)
where n is number of valence electrons to be transferred, F96.5kC/mol is Faraday constant, and E is voltage added or half-cell potential.
G governs the direction and equilibrium of a chemical reaction.
If G < 0, free energy decreases, so reaction will be proceeded easily.
Nernst equation is a log relationship derived from basic thermodynamic equations, e.g.,
G = -RTlnK (4)
where R=8.31J/mol·K is gas constant.
Combining Eq.(3) and Eq.(4), we can get the generalized term:
E = (RT/nF)lnK (5)

## Slide 13: Transducers: Electrodes for Biosensors
A two half–cell loops:
2 ionic solutions w/ different concentrations are separated by a ion-selective bridge, the measured potential can be derived from Nernst Eq.(5):
where a1, a2 are ionic activity . In low concentration,
a1=[a1], a2=[a2].
Concentration cells:
A standard half-cell potential E is determined with a1=1 (1M). If the activity is changed from 1, it will change and can be given by Nernst equation as
General term for oxidation-reduction reaction:
The Nernst equation will be
Calculate deviation from E0

## Slide 14: Transducers: Electrodes for Biosensors
4. Equivalent circuit of electrode-electrolyte interface:
An electrode is a very complicated nonlinear circuit (why?).
E
Rd
Cd
Rs
E
Rd
Cd
Rs
E
Rs
E
Rd+Rs
HF: 1/c<<Rd
IF:
LF: 1/c>>Rd
Freq. characteristics:
Rd: leaky resistor at interface (double layer);
Cd: concentration related capacitor at double layer;
Rs: resistor in the electrolyte.
LF
IF
HF

## Slide 15: Transducers: Electrodes for Biosensors
Skin-electrode model:
electrode
Assume Rgel=0
Epidermis: behaves like a membrane
for ion transfer w/:
Eskin, Rep and Cep.
The epidermis impedance
is strongly f-dependent, e.g.,
varying from 200k at
1Hz to 200  at 1MHz for
1 cm2 electrode.
The major role is the stratum
corneum.
Selective to Na+, K+
and Cl- channels.
Mostly neglected.
Dermis: behaves like a resistive
layer w/ RD.
Usually, the stratum corneum is removed
by mechanical abrasion or chemical way to
improve conductivity – critical to reduce
noise. In this case, Eskin=0, Cep=0, and
Rs= Rep+ RG+ RD.
Also, if the electrode is moved w/ respect to electrolyte, local ionic charge distribution will be disturbed, resulting in a momentary change in Eskin,. Because a nonpolarizable electrode, e.g., Ag/AgCl has low overpotential E0, this does not pose a problem. However, in precise sensing, motion-artifact btw gel and skin Eskin, may cause a major artifact. This can be removed by removal of stratum corneum. Side effect!